## Simple Mask
Vendor
Model
Product | Spec category | Specs, WHO (simple mask derived) | Meet spec? | Comments
Simple oxygen mask | Technical | Mask is soft, transparent, well-fitting moulded, with two valve-less side vents to allow for entrainment of ambient air and passage of exhalation.
o Multiple sizes for children and adults
Nose clip is soft, malleable and adjustable.
Rubber or soft plastic:
o Transparent PVC or other equivalent material,
o FDA Title 21 / USP plastic class VI compliant, certified for medical use
o Semi-rigid, allowing freedom of movement: hardness> 60 Shore A (ASTM D-2240)
Tubing:
o Low resistance tubing, round in section
o Anti-kinking tubing
o 3–5 mm internal diameter and 7–8 mm external diameter
o compatible for use with oxygen
o 1.5-2 m tubing length
o with “standard” and “universal” hose end connector
Non-sterile; single use.
Individually packed.
| Warranty | Minimum 3 year shelf-life, recommend 5
| QMS | ISO 13485 (medical device QMS)
ISO 9001 (General QMS, optional)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
| Product performance standards | ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment.
ISO/DIS 17256 Anaesthetic and respiratory equipment – Respiratory therapy tubing and connectors.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
The words "for single use".
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Venturi Mask
Vendor
Model
Product | Spec category | Specs, WHO (simple mask derived) | Meet spec? | Comments
Venturi mask | Technical | Air-entrainment mask (with percent O2 lock) that delivers oxygen with a specific concentrations from 24–60% minimum.
The kit of the mask includes:
o Tubing
o humidity cup
o multiple jets or "diluters" (min XX, up to 6), which are colour coded clearly indicating percentage of oxygen delivered.
· blue = 2-4 L/min = 24%; white = 4-6 L/min = 28%; orange = 6-8 L/min = 31%; yellow = 8-10 L/min = 35%; red = 10-12 L/min = 40%; green = 12-15 L/min = 60%.
Mask is soft, transparent, well-fitting moulded, with fittings for oxygen entrainment jet/diluter, and two valve-less side vents to allow for entrainment of ambient air and passage of exhalation.
Nose clip is soft, malleable and adjustable.
Rubber or soft plastic:
o Transparent PVC or other equivalent material,
o FDA Title 21 / USP plastic class VI compliant, certified for medical use
o Semi-rigid, allowing freedom of movement: hardness> 60 Shore A (ASTM D-2240)
Tubing:
o Low resistance tubing, round in section
o Anti-kinking tubing
o 3–5 mm internal diameter and 7–8 mm external diameter
o compatible for use with oxygen
o 1.5-2 m tubing length
o with “standard” and “universal” hose end connector
Non-sterile; single use.
Individually packed.
| Warranty | Minimum 3 year shelf-life, recommend 5
| QMS | ISO 13485 (medical device QMS)
ISO 9001 (General QMS)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
| Product performance standards | ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment.
ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications Part 1: Evaluation and testing within a risk management process.
ISO/DIS 17256 Anaesthetic and respiratory equipment – Respiratory therapy tubing and connectors.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
The words "for single use".
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Nasal Cannulae
Vendor
Model
Product | Spec category | Specifications adapted from WHO-UNICEF (2019) | Meet spec? | Relevant certificate provided? | Comments
Nasal cannulae | Description | A non-sterile, semi-rigid tube with nasal prongs designed to be inserted into the nostrils of a patient, and held in place with a headstrap, to administer oxygen (O2). It is commonly known as "nasal prongs". This is a single-use device. (Source: GMDN)
| Technical | Oxygen cannulae are pastic tubes shaped as two prongs delivering an air/oxygen mixture into the nasal cavities and are connected with an oxygen administration circuit.
o Straight, curved and flared tip prongs.
o Multiple sizes for neonates, infants, children and adult (spacing, diameter, length, flow capacity).
o Tubing and prong components packaged as ready to use circuits, complete with adapters to connect to oxygen therapy circuits
Rubber or soft plastic:
o Transparent PVC or other equivalent material,
o FDA Title 21 / USP plastic class VI compliant, certified for medical use
o Semi-rigid, allowing freedom of movement: hardness> 60 Shore A (ASTM D-2240)
o Standard, latex free and phthalate/DEHP-free versions, available for tubes and prongs.
Tubing:
o Low resistance tubing, round in section
o Anti-kinking tubing (star lumen or other design)
o 3–5 mm internal diameter and 7–8 mm external diameter
o compatible for use with oxygen
o 1.5-2 m tubing length
o with “standard” and “universal” hose end connector
o Tie sliding and over-the-ear tubing adjustable options, with or without ear guards.
Non-sterile; single use.
Suitable for cleaning and disinfection
Individually packed.
Environmental:
o Storage: ambient temperature of at least 5–50 °C, relative humidity of at least 15–95% non condensing.
o Use: ambient temperature of at least 5–45 °C, relative humidity of at least 15–90% non condensing
| Warranty | Minimum 3 year shelf-life, recommend 5
| QMS | ISO 13485 (medical device QMS)
ISO 14971 (application of risk management)
ISO 10993-1 Biological evaluation of medical devices - part 1: Evaluaiton and testing within a risk management process
| Regulatory | CE
FDA
| Product performance standards | ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
The words "for single use".
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Non-rebreather masks
Vendor
Model
Product | Spec category | Specs, WHO (simple mask derived) | Meet spec?
Non-rebreather mask with reservoir bag | Technical | Two unidirectional valves, one that closes during inspiration to prevent room air mixing with oxygen in a reservoir bag; and one that closes during exhalation to prevent exhaled respiratory gases from entering the reservoir bag.
Mask is soft, transparent, well-fitting moulded, with two side vents.
Nose clip is soft, malleable and adjustable.
Rubber or soft plastic:
o Transparent PVC or other equivalent material,
o FDA Title 21 / USP plastic class VI compliant, certified for medical use
o Semi-rigid, allowing freedom of movement: hardness> 60 Shore A (ASTM D-2240)
Tubing:
o Low resistance tubing, round in section
o Anti-kinking tubing
o 3–5 mm internal diameter and 7–8 mm external diameter
o compatible for use with oxygen
o 1.5-2 m tubing length
o with “standard” and “universal” hose end connector
15/22 mm diameter ventilation tubing
Non-sterile; single use.
Individually packed.
Sizes: adult and pediatric
| Warranty | Minimum 3 year shelf-life, recommend 5
| QMS | ISO 13485 (medical device QMS)
ISO 9001 (General QMS)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
| Product performance standards | ISO 15001 Anaesthetic and respiratory equipment – Compatibility with oxygen.
ISO 18190 Anaesthetic and respiratory equipment – General requirements for airways and related equipment.
ISO 18562-1 Biocompatibility evaluation of breathing gas pathways in healthcare applications Part 1: Evaluation and testing within a risk management process.
ISO/DIS 17256 Anaesthetic and respiratory equipment – Respiratory therapy tubing and connectors.
ISO 15223-1 Medical devices – Symbols to be used with medical device labels, labelling and information to be supplied – Part 1: General requirements.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Expiry date by year and month, prefixed by the word "EXP" (or equivalent harmonised symbol).
The words "for single use".
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## RFI_MFLab (hidden)
| 关于产品 | [9 empty] | 关于贵司
产品类型 Product Type | 产品标准要求 Product Specifications | 产品型号（请列出贵公司符合左侧标准要求的产品，并邮件提供电子版产品用户手册） Product Model (s) | 产品排期（假设下周下订单，预计能够供的最早出货排期至何时（日期）？ Lead time if order placed next week (date) | 本次订单所需数量 Units Required | 公司当前可出售的数量 Units Available | 法规批准/认证Regulatory Approval/Certification |  | 产品性能标准Standards, for product performance |  | 产品单价（美金） Unit price (USD) | 制造商标准 Standards, for manufacturer |  | 预付款要求 Pre-payment requirement | 公司以前有向埃塞俄比亚出口的经验吗？Does company have previous export experience to Ethiopia? | 图片示例 Picture
[6 empty] | 标准 | 如果有，请填写 “X” | 标准 | 如果有，清填写 “X” |  | 标准 | 如果有，清填写 “X”
简单氧气气面罩 Simple Face Masks | 成人式： 医用级 | LB1110, Adult(LA), 5.5mm x 2m绿色管 |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X | 生产前付清 | 有
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
不重吸面罩 Oxygen Mask with reservoir bag | 成人式： 氧气浓度：支持80%至100％的氧气浓度 管长: 2m 使用标准接头；管路（氧气管线）不弯曲 材质：聚氯乙烯（PVC）； 透明；医用级 | LB1310, Adult(LA), 5.5mm x 2m绿色管 氧袋1000ml |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
文丘里面罩（氧浓度可调面罩） Venturi Mask | 成人式：能以8至10 L / min 的流速输送低至中等流量的氧气，特定氧浓度为30％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,31%(氧气流速6L/min±10%),白色，其他符合 |  | 300 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
| 成人式：能以10至12 L / min 的流速输送低至中等流量的氧气，特定氧浓度为40％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,40%(氧气流速8L/min±10%),粉色，其他符合 |  | 300 |  | CE | X | ISO 18190
[6 empty] | FDA | X | ISO18562-1 | X
[6 empty] | 出口许可证 | X | ISO/DIS 23368
| 成人式：能以12至15 L / min 的流速输送低至中等流量的氧气，特定氧浓度为60％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,60%(氧气流速8L/min±10%),红色，其他符合 |  | 300 |  | CE | X | ISO/DIS 17256 | X
[6 empty] | FDA | X | ISO 15223-1 | X
[6 empty] | 出口许可证 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 中等大小 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带, M |  | 130 |  | CE |  | ISO 18562-1:2017 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 17510:2015 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X |  |  |  | ISO14971 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 大码。 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带，L |  | 130
气管导管 Endotracheal Tube | 管： 尺码：7。 内径：7.0mm。 外径：9.3mm。 长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 袖口： 直径：24mm | LB5010C, 带囊，ID 7.0mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X |  | ISO14971 | X
| 管： 尺码：7.5 内径：7.5mm。长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 | LB5010C, 带囊，ID 7.5mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X