## Pulse Oximeter, HAND-HELD
Vendor
Model
Product | Spec category | Specifications adapted from WHO-UNICEF 2019 | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
Pulse oximeter, handheld | DESCRIPTION | A portable, battery-powered, photoelectric device intended for the transcutaneous measurement and display of haemoglobin oxygen saturation (SpO2). The signals, typically produced by light-emitting diodes (LEDs) and a receiving detector in a probe, or directly built-in, are used to make the measurements using the principle of spectrophotometry. The oximeter displays the SpO2 values and may calculate/display other parameters, e.g. pulse rate, electrocardiogram (ECG). The device is typically applied to the fingertip or around the wrist; it may be used by health care facilities, emergency services, or in the home. (Source: GMDN)
Continuously measures and displays main parameters: %SpO2, pulse rate, plethysmographic waveform (and possibly other indicators of signal quality), alarm messages, battery state indication.
Suitable for use with adults, paediatrics, and neonates (with appropriate probes)
Suitable for all skin pigmentations.
Suitable for detection in low perfusion conditions (as per ISO 80601-2-61).
Display must allow easy viewing in all ambient light levels.
SpO2 detection to include the range: 70–100%.
SpO2 resolution: 1% or less.
SpO2 accuracy (in the range at least 70–100%): within ± 2% under ideal conditions of use, and within ± 3% for all patients and perfusion/movement conditions.
Pulse rate detection to include the range: 30–240 bpm.
Pulse rate resolution: 1 bpm or less.
Pulse rate accuracy within ± 3 bpm.
Data update period for valid data displayed ≤ 10 s.
Data storage (optional):
Internal data storage for patient trends and event log (optional).
Data interface, suitable for exporting data to external software (optional)
Automatic correction for movement and ambient light artefacts (as per ISO 80601-2-61.
Alarms:
Audible and visual alarms for low/high saturation and pulse rate, threshold set by user. (adjustable)
Audible and visual alarms for sensor error or disconnected, system errors, low battery.
Alarm override and temporary silencing function.
Probes:
Capable of working with adult, paediatric and neonatal reusable probes.
Supplied with, adult, paediatric and neonatal reusable probes.
Suitable for cleaning and decontamination (probes and unit)
Device + probe to weigh < 400 grams
Power supply:
Operates by replaceable battery power supply, either rechargeable or single use.
External or built-in AC battery charger, if rechargeable type.
Plug style as per local supply.
Charger, if used, to have protection against over-voltage and over-current line conditions and be certified to IEC 60601-1.
Protection against defibrillator discharges and electrosurgical units.
Running time on battery only ≥ 12 hours.
Documentation:
User and maintenance manuals, hard and soft copies, to be supplied in local language, or agreed other language if local not available. English version must also be available.
Certificate of calibration and inspection to be provided.
| Environmental | Enclosure to have ingress protection level IPX2 or better.
Functionality in demanding environments as per tests in ISO 80601-2-61 e.g., shock, vibration, and IEC 60068-2-31 e.g., free-fall
Storage conditions: ambient temperature of at least 5–50 °C, relative humidity of at least 10–95% non condensing.
Operating conditiongs: ambient temperature of at least 5–40 °C, relative humidity of at least 10–90% non-condensing.
| Warranty | 2 years recommended, 1 year mandatory
| QMS | ISO 13485 (medical device QMS)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
| Product performance & safety standards | ISO 80601-2-61
IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance.
IEC 60601-1-1 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems.
IEC 60601-1-2 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests.
IEC 60068-2-31 Environmental testing – Part 2-31: Tests – Test Ec: Rough handling shocks, primarily for equipment-type specimens.
IEC 62133 – Secondary cells and batteries containing alkaline or other non-acid electrolytes – Safety requirements for portable sealed secondary cells. Part 1: Nickel, Part 2: Lithium.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Pulse Oximeter, FINGERTIP
Vendor
Model
Product | Spec category | Specifications adapted from WHO-UNICEF 2019 | Meet requirement? (Y/N) | Type of supporting documents (select from dropdown) | Page no. in supporting documents | Comments
Pulse oximeter, handheld | DESCRIPTION | A portable, battery-powered, photoelectric device intended for the transcutaneous measurement and display of haemoglobin oxygen saturation (SpO2). The signals, typically produced by light-emitting diodes (LEDs) and a receiving detector in a probe, or directly built-in, are used to make the measurements using the principle of spectrophotometry. The oximeter displays the SpO2 values and may calculate/display pulse rate. The device is typically applied to the fingertip; it may be used by health care facilities, emergency services, or in the home. (Adapted from: GMDN)
Measures and displays main parameters %SpO2, pulse rate, and plethysmographic wave/pulse (and possibly other indicators of signal quality), battery state indication.
Intended use-case: spot-check
Suitable for use with adults
CAN be used for paediatrics if sized appropriately, and ONLY for neonates with appropriate external probes connected
Suitable for all skin pigmentations.
Suitable for detection in low perfusion conditions (as per ISO 80601-2-61, test method to be described)
SpO2 detection to include the range: 70–99%.
SpO2 resolution: 1% or less.
SpO2 accuracy (in the range at least 70–100%): within ± 3% under ideal conditions of use and for all patients and perfusion/movement conditions.
Pulse rate detection to include the range: 30–240 bpm.
Pulse rate resolution: 1 bpm or less.
Pulse rate accuracy within ± 3 bpm.
Data update period for valid data displayed ≤ 10 s.
Automatic correction for movement and ambient light artefacts as per ISO 80601-2-61.
Suitable for cleaning and decontamination
Display:
Must allow easy viewing in all ambient light levels.
Plethysmographic visualization, preferabley waveform but pulse acceptable
Battery status
Sensor error / signal fault
Power supply:
Batteries may be single use, or rechargeable (preferred) with external AC battery charger, or by USB connection.
Charger, if used, to have protection against over-voltage and over-current line conditions and be certified to IEC 60601-1.
If rechargeable batteries are used:
If expected lifetime of unit is more than expected lifetime of the batteries, rechargeable batteries must be replaceable by the user.
Hours of continuous use, or number of tests, per battery set shall be stated
Batteries must allow at least 2500 spot checks calculated at 30 s per spot check, or at least 21 hours of operation.
Automatic power-off.
Documentation:
User and maintenance manuals, hard and soft copies, to be supplied in local language, or agreed other language if local not available. English version must also be available.
Certificate of calibration and inspection to be provided.
| Environmental | Enclosure to have ingress protection level IPX2 or better.
Functionality in demanding environments as per tests in ISO 80601-2-61 e.g., shock, vibration, and IEC 60068-2-31 e.g., free-fall
Storage conditions: ambient temperature of at least 5–50 °C, relative humidity of at least 10–95% non condensing.
Operating conditiongs: ambient temperature of at least 5–40 °C, relative humidity of at least 10–90% non-condensing.
Specific requirements for altitude may be required, depending on use site
Display must allow easy viewing in all ambient light levels.
| Warranty | 2 years recommended, 1 year mandatory
| QMS | ISO 13485 (medical device QMS)
ISO 14971 (application of risk management)
| Regulatory | CE
FDA
Japan, Canada, Australia
| Product performance & safety standards | ISO 80601-2-61
IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance.
IEC 60601-1-1 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems.
IEC 60601-1-2 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests.
IEC 60068-2-31 Environmental testing – Part 2-31: Tests – Test Ec: Rough handling shocks, primarily for equipment-type specimens.
IEC 62133 – Secondary cells and batteries containing alkaline or other non-acid electrolytes – Safety requirements for portable sealed secondary cells. Part 1: Nickel, Part 2: Lithium.
| Packaging | Name and/or trade mark and address of the manufacturer.
Product name.
Product reference.
Type of product and main characteristics.
Performance testing information against the mentioned standards.
Lot number prefixed by the word "LOT" (or equivalent harmonised symbol).
Information for particular storage conditions (temperature, pressure, light, humidity, etc.), as appropriate (or equivalent harmonised symbol), if applicable.
Information for handling, if applicable (or equivalent harmonised symbol).
If the packaging is not transparent, it must bear a diagram (preferably actual size) showing the essential parts of the product and indicating the position of the product in the packaging.
Gross Weight.
Cubic Measurement.
All indicated at least in English.

## Patient Monitor=
Vendor name:
Model details:
Product | Spec category | Specs, WHO-UNICEF derived | Meets requirement? (Yes/No/NA) | Type of supporting document (select from drop down) | Page No. | Comments
Patient monitor multiparametric: moderate | Product description | Patient monitor multiparametric (moderate) – for ECG, non-invasive blood pressure (NIBP), oxygen saturation (SpO2), respiratory rate (RR) and temperature (TEMP). The monitors are designed to give information on one screen and to provide multiple information that is needed to understand the patient condition. It has emerged as a monitor to offer flexible solution for varying critical care need.
| Technical | Continuous display on the screen of the following parameters: ECG, respiration rate and heart rate, non-invasive blood pressure, body temperature and SpO2
Dynamic digital display that can show all active parameters
Unwanted parameters can be deselected from display
Operator can set audio-visual alarm levels for low or high levels of each parameter independently
Operates from mains voltage or from internal rechargeable battery
ECG patient connectors that are sterilizable and reusable are preferred
Trend display of each parameter to fall within ranges and accuracy where indicated:
ECG: - Minimum 6-lead ECG measurement and selectable display - patient connectors that are sterilizable and reusable are preferred - trace signal velocity of at least 25 mm/s
Heart rate: measurement range to be at least 30–250 bpm, with accuracy better than ± 5 bpm and minimum gradation 1 bpm
SpO2: measurement range at least 70–99 %, with accuracy better than ± 3% and minimum gradation 1%
NIBP: - Blood pressure monitoring range at least 30–270 mmHg, minimum gradation 1 mmHg. Cuff sizes: paediatric and adult. User selectable measurement intervals - Internal pump for cuff inflation for non-invasive blood pressure measurement, with over pressure protection
Temperature: - Range at least 30–40 °C, minimum gradation 0.1 °C - Temperature probe to be reusable, external skin contact type, but consumable to protect between patients or disinfection method explained
Respiration rate: measurement range at least 0–100 bpm, minimum gradation 1 bpm
Automatic and programmable memory
Storage of continuous monitoring data
LCD or TFT screen with:
analogue shape signals and numerical values visualization
settable limits for the measured variables
Design must enable use in demanding environments, e.g. shock, vibration and free fall tests as per tests
Protections of all the functions against defibrillator discharges and electrosurgical units
Pace-maker detection
Data management functions (preferable)
Capable of operating continuously in ambient temperature of 10–40 °C and relative humidity of 15–85% (90% preferable)
Enclosure to have ingress protection level IPX1 or better
| User adjustable settings and alarms | Alarm override and temporary silence facility to be included
Audio-visual alarms required: high and low levels for each parameter (operator variable settings), sensor/wire/probe disconnected, low battery, cuff leak, cuff, disconnect, hose leak, inflation/deflation errors, failure to take successful reading, low battery notice
Power failure alarm
| Consumables | ECG electrodes (if applicable): 100 sets
| Accessories | All the cables, sensors and connectors needed for full monitor functionality are to be included in the bid
Lead ECG cable: 2 per equipment
Sets of ECG connection electrodes (if reusable type): 5 sets
Tubes electrode gel (if required): 5 tubes
Reusable SpO2 probes adult: 2 probes
Reusable SpO2 probes paediatric use: 2 probes
Blood pressure – non-invasive: 3 paediatric reusable cuffs; 3 adult reusable cuffs
External skin temperature probes: 2 probes
Battery: 1 set
| Spare parts | 1-year spare parts kit as per preventive maintenance program, including but not exclusively: sets of spare fuses (if non-resettable fuses used) and battery
| Power supply voltage, frequency and plug (may vary across countries) | Operated by line electrical power supply with internal replaceable rechargeable battery backup allows operation for at least 1 hour in the event of power failure
Operates from AC power electric line: 100–240 V~/50–60 Hz
Main power cable and plug adapted for various countries
Mains power cable length ≥ 2.5 m
Protections against over-voltage and over-current line conditions
Protection against defibrillator discharges and electrosurgical units
Automatic switch between battery and mains powered modes, when recharging or in mains power failure
The display shows which power source is in use
Compliance with electrical standards and regulations
| Warranty | 5 years for efficiency and quality of the product (software upgrades included)
Ensure availability of accessories and spare parts for at least 5 years
| Manufacturing Standards | ISO 13485 - Certified quality management system for medical devices
ISO 9001 - General quality management
| Regulatory | Free sales certificate (FSC) or certificate for exportation of medical device provided by the authority in manufacturing country
CE under EU's MDR directive, or
FDA, or
Proof of regulatory compliance as per the product’s risk classification of other IMDRF founding member country: Canada, Australia, Japan
| Product performance standards | IEC 60601-1 Medical electrical equipment – Part 1: General requirements for basic safety and essential performance
IEC 60601-1-1 Medical electrical equipment – Part 1-1: General requirements for safety – Collateral standard: Safety requirements for medical electrical systems
IEC 60601-1-2 Medical electrical equipment – Part 1-2: General requirements for basic safety and essential performance – Collateral standard: Electromagnetic compatibility – Requirements and tests
IEC 60601-1-8 General requirements for basic safety and essential performance – Collateral standard: General requirements, tests and guidance for alarm systems in medical electrical equipment and medical electrical systems
IEC 80601-2-49 Medical electrical equipment – Part 2-49: Particular requirements for the basic safety and essential performance of multifunction patient monitoring equipment
IEC 80601-2-30 Particular requirements for the basic safety and essential performance of automated non-invasive sphygmomanometer
ISO 80601-2-61 Medical electrical equipment – Part 2-61: Particular requirements for basic safety and essential performance of pulse oximeter equipment
IEC 60601-2-27 Particular requirements for the basic safety and essential performance of electrocardiographic monitoring equipment
Preferable if tested for:
IEC 62366-1 Medical devices – Part 1: Application of usability engineering to medical devices
IEC 60068-1:2013 Environmental testing – Part 1: General and guidance
IEC 60068-2-31 Environmental testing – Part 2-31: Tests: Rough handling shocks, primarily for equipment-type specimens
| Packaging and documentation requirements (French language) | Name and/or trademark of the manufacturer
Electrical power input requirements (voltage, frequency and socket type) and safety use and storage (keep away from oil, grease and petroleum-based or flammable products as well as smoking or open flames)
Model or product reference
Information for particular storage conditions (temperature, pressure, light, humidity)
Set: user and maintenance manuals, hard and soft copies, in English (mandatory) and local languages (preferable)
Certificate of calibration and inspection
Troubleshooting, calibration and routine maintenance
List of all spare parts and accessories, with part numbers and contact details for parts supply
Document with contact details of manufacturer, supplier and local service agent

## RFI_MFLab (hidden)
| 关于产品 | [9 empty] | 关于贵司
产品类型 Product Type | 产品标准要求 Product Specifications | 产品型号（请列出贵公司符合左侧标准要求的产品，并邮件提供电子版产品用户手册） Product Model (s) | 产品排期（假设下周下订单，预计能够供的最早出货排期至何时（日期）？ Lead time if order placed next week (date) | 本次订单所需数量 Units Required | 公司当前可出售的数量 Units Available | 法规批准/认证Regulatory Approval/Certification |  | 产品性能标准Standards, for product performance |  | 产品单价（美金） Unit price (USD) | 制造商标准 Standards, for manufacturer |  | 预付款要求 Pre-payment requirement | 公司以前有向埃塞俄比亚出口的经验吗？Does company have previous export experience to Ethiopia? | 图片示例 Picture
[6 empty] | 标准 | 如果有，请填写 “X” | 标准 | 如果有，清填写 “X” |  | 标准 | 如果有，清填写 “X”
简单氧气气面罩 Simple Face Masks | 成人式： 医用级 | LB1110, Adult(LA), 5.5mm x 2m绿色管 |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X | 生产前付清 | 有
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
不重吸面罩 Oxygen Mask with reservoir bag | 成人式： 氧气浓度：支持80%至100％的氧气浓度 管长: 2m 使用标准接头；管路（氧气管线）不弯曲 材质：聚氯乙烯（PVC）； 透明；医用级 | LB1310, Adult(LA), 5.5mm x 2m绿色管 氧袋1000ml |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
文丘里面罩（氧浓度可调面罩） Venturi Mask | 成人式：能以8至10 L / min 的流速输送低至中等流量的氧气，特定氧浓度为30％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,31%(氧气流速6L/min±10%),白色，其他符合 |  | 300 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
| 成人式：能以10至12 L / min 的流速输送低至中等流量的氧气，特定氧浓度为40％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,40%(氧气流速8L/min±10%),粉色，其他符合 |  | 300 |  | CE | X | ISO 18190
[6 empty] | FDA | X | ISO18562-1 | X
[6 empty] | 出口许可证 | X | ISO/DIS 23368
| 成人式：能以12至15 L / min 的流速输送低至中等流量的氧气，特定氧浓度为60％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,60%(氧气流速8L/min±10%),红色，其他符合 |  | 300 |  | CE | X | ISO/DIS 17256 | X
[6 empty] | FDA | X | ISO 15223-1 | X
[6 empty] | 出口许可证 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 中等大小 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带, M |  | 130 |  | CE |  | ISO 18562-1:2017 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 17510:2015 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X |  |  |  | ISO14971 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 大码。 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带，L |  | 130
气管导管 Endotracheal Tube | 管： 尺码：7。 内径：7.0mm。 外径：9.3mm。 长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 袖口： 直径：24mm | LB5010C, 带囊，ID 7.0mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X |  | ISO14971 | X
| 管： 尺码：7.5 内径：7.5mm。长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 | LB5010C, 带囊，ID 7.5mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X