## QA
Request for Quotations (RFQ) for Diagnosis and Monitoring Support devices - Q&A
RFQ/CHAI/DMS/0122
No. | Clarification requested | Answer
1 | Requirement in the RFQ is 3 years,But in Annex A the warranty for Pulse Oximeter, HAND-HELD is 2 years recommended, 1 year mandatory,for Pulse Oximeter, FINGERTIP is 2 years recommended, 1 year mandatory.for Patient Monitor is 5 years for efficiency and quality of the product.Could you please kindly advise which warranty is the right one | Desired warranty is as detailled in Annex B for each product
2 | Do we need to install the product on site or on line guide is enough,do we need to include the on-site installtion in our quotation. Within the warranty period,are there any detailed requirement for local service.Do we need to provide maintaince and repair on site or on line. | - On-line guides for installation are acceptable. - Within the warranty period, spares for repair, as well as consumables for operations, shall be easily and rapidly available. Further, any repairs required under warranty are the responsibility of the vendor. Whether this is in-person or done virtually is dependent on need and/or requirements to resolve specific issue.
3 | In the RFQ,The vendor will finish the obligation and get the payment after the supply and inspection.The inspection will be carried out in the delivery address or the end user's hospital?Could we request for a time limit for the inspection after the delivery. | Inspections will be carried out centrally, after the goods have cleared customs.
4 | We will have a third party Logistics company to help on the feight,and a local service provider help on the service.Could we provide these two companies information separately in the Vendor information in Annex-A. | Yes
6 | As a wholesaler,could we quote 2 models from different manufactures for one product. | No please quote your most prefered product as each supplier will be evaluated based on one product
7 | Pulse Oximeter,FINGERTIP and HANDHELD :The parameters of the oximeter are required to be above IPX2 level. While CE registration and FDA registration have no strict requirements for IP level. Ipx2 only increases the application scenario of out of hospital transfer compared with ipx1. Please kindly confirm whether ipx1 is acceptable? | Indeed, neither CE nor FDA indicate IP rating. It is the manufacturer to claim and provide evidence of product rating when submitting their dossier for regulatory approval. As per WHO-UNICEF specifications, we are requesting devices with IPX2 rating. If this is not available (or covered under your regulatory approval), we do still encourage you to submit your offer. Kindly provide any details or comments in Annex B.
8 | Pulse Oximeter, FINGERTIP“Batteries may be single use, or rechargeable (preferred) with external AC battery charger, or by USB connection”, Please kindly confirm could we quote for two alternative config for different battery | Yes. Please clearly indicate differences in quotations
9 | Can we present more than one model for the items? | No please quote your most prefered product as each supplier will be evaluated based on one product
10 | Should we provide the patient monitor with a trolley? | No, trolleys have not been specified for the patient monitors.
11 | It has been indicated that spare parts are needed for the patient monitor. Is it mandatory to quote or optional? | It is not mandatory, but this will be reflected in the technical evaluation of your offer.