## Instructions for specs VIE tank
Please read the instructions carefully to ensure that all required documentation is submitted.
Only complete bids with all tabs filled and supplemental documentation attached will be considered.
Tab: "L2G"
This table contains general specifications for the the liquid to gas filling equipment included in the RFP.
Fill column C to indicate whether the product complies with the specification.
Fill column D to indicate what type of supporting documentation is being provided as evidence of compliance with the specification.
Fill column E with comments as applicable. If "other" is selected in Column E, please specify the type of document.
Tab: "Tanks (General)"
This table contains general specifications for the bulk storage tanks included in the RFP.
Fill column C to indicate whether the product complies with the specification.
Fill column D to indicate what type of supporting documentation is being provided as evidence of compliance with the specification.
Fill column E with comments as applicable. If "other" is selected in Column E, please specify the type of document.
Tab: "Tanks (Product Specific)"
This table product contains specifications for bulk storage tank product line submitted by vendor in RFP.
Vendors must fill rows 3-8 for product line included in proposal
Tab: " L2G Systems installation and commissioning"
This table contains the installation and commissioning requirements for the L2G system included in the RFP.
The installation and commisionning have to follow the tank and vaporiser manufacturer recommendations/requirements
The installation and commisionning will be subjected to a third party authorized inspection agency ( AIA )
Please review these requirements carefully before preparing the technical proposal.
Vendors must submit proposals indicating that these requirements are met.
Tab: " First fill LOX manufacturer"
This table contains the installation and commissioning requirements for the First fill LOX manufacturer part of the commissionning included in the RFP.
The installation and commisionning have to follow the tank and vaporiser manufacturer recommendations/requirements
The installation and commisionning will be subjected to a third party authorized inspection agency ( AIA )
Please review these requirements carefully before preparing the technical proposal.
Vendors must submit proposals indicating that these requirements are met.
Tab: " First fill LOX transport"
This table contains the installation and commissioning requirements for the First fill LOX transport part of the commissionning included in the RFP.
The installation and commisionning have to follow the tank and vaporiser manufacturer recommendations/requirements
The installation and commisionning will be subjected to a third party authorized inspection agency ( AIA )
Please review these requirements carefully before preparing the technical proposal.
Vendors must submit proposals indicating that these requirements are met.
Query period: | September - October 2023
Query email address: | drcprocurement@clintonhealthaccess.org

## L2G
Component | Specification | Meet requirement? | Type of supporting documents | Comments
Overview | Output operating with a cryopump and high-pressure vaporizer to fill compressed gas cylinders for medical oxygen service
Capacity to fill 20 J size cylinders per hour
Average run time: 8 hours per day
Fill pressure: 200 bar
Provision and installation of liquid to gas filling station equipment shall be in accordance with the approved vendor proposed layout
Configuration | The liquid to gas (L2G) filling station consists of the following components:
• A 20 tonnes bulk storage tank (see tanks(General) tab)
→ Interconnected configuration, must be reflected in P&ID
• Cryogenic liquid oxygen (LOX) pump
• High-pressure vaporizer
• Medical grade copper pipes, fittings, and accessories
→ Liquid lines are rated for cryogenic service
• Pressure reducing system
• Filling manifold complete with:
→ 2 rows of cylinder connection points ( 40 connection points)
→ Bullnose valve connection (Compatible with BS 341 cylinder valve outlets)
→ Secure fastening for cylinders (e.g., chain, bracket)
→ Protection cage
• Purge compressor/vacuum to empty cylinders
Cryogenic LOX Pump | Oil lubricated reciprocating pump ( Please specify the make and model)
Maximum Output pressure: 200 bar
≥ Duplex configuration
Vacuum jacket
Mechanism for degassing
Feed and return lines
Safety guard
Suction filter
Relief valve
Safety check / non-return valves
Low pressure seals
High-Pressure | High pressure vaporizer( Please specify the make and model)
• Rated capacity: 150 m3/hr
• Stainless steel lined
• Rated working pressure: ≥ 200 barg
Design /maximum allowable working temp -196°C /+50°C
Inlet and outlet flange availability: ASME B16.5, BS PR C1, EN 1092, M40x2 thread, DIN 2635
All models designed according to PED 2014/68/EU, ASME SEC.VIII DIV.I, or ANSI 831.3
System Operations | Program logic-controlled System
Touch-screen control panel in English or French
Integrated O2 purity analyzer and monitor for medical application, paramagnetic type (or indicate otherwise)
Displayed parameters:
• Purity Oxygen concentration [%], with +/- 1% accuracy)
• System Leakage
• Pressure indication in line and headers
• Temperature
• Overall system status and time to next service
• Cummulative operating hours
Emergency shutoff
High discharge pressure switch
Loss of prime / cavitation shutdown system
Gas leakage detection
Remote monitoring feature (optional)
Power requirements | Cryopump and high-pressure vaporizer operations: 380 volts /50 Hz (3 Phase)
Control system operations: 220 VAC / 50 Hz (Single phase)
Surge protection/suppression systems
Power meter to measure the electricity consumption of the filling station
Spares | Three (3) year spare parts kits as per recommended preventative maintenance programme clearly defined in a disaggregated list comprising part numbers, descriptions, and unit cost as well as indicating brand/model specifics by the manufacturer
Installation, testing and commissioning | Detailed description of any relevant experience of importing/suppling liquid to gas filling station equipment
Evidence of effectively managed after sales service, including but not limited to installation, testing, and commissioning, repair, warranties honored, and provision of spare parts
Pre-shipment: Certificate of quality, calibration, and inspection
• Provision of medical LOX supply for operational first fill during filling system commissioning at CUK (see first fill LOX manufacturer and first fill LOX transport tabs)
On-site:
• Installation to include training on start-up, use, and preventive/curative maintenance
• Inspection, testing and commissioning should be completed
Third-party inspection:
• Conformity of installation shall be verified by Authorized Inspection Authority (AIA)
Documentation | P&ID of proposed equipment configuration (Fully labelled with Element Discription/Notes and Legend)
Hard and soft copies, in English or French of the following:
• Certificate of quality, calibration, and inspection
User manual, detailing:
• Protocols for operations and filling instructions
• List of procedures and equipment required for cleaning
• List of procedures and equipment for preventive maintenance
• Troubleshooting chart
Contact details of manufacturer, and authorized distributors (if applicable), and local service agents
Product and model reference (brochure)
Warranty | All components of the L2G cylinder filling station should have at least 3 years’ warranty period after commissioning that covers replacement parts and labor for any manufacturer defects
The supplier must ensure the availability of spare parts for the duration of 8 years and indicate lead times
Regulatory | PED 2014/68 EU (indicate NB) Or ASME “U” stamp
QMS | ISO 9001: Quality Management Systems - Requirements
Process or component specific performance standards (or equivalent) | EIGA 159/21: Reciprocating Cryogenic Pumps and Pump Installatoins for Oxgyen, Argon, and Nitrogen
EIGA Doc 33: Cleaning of Equipment for Oxygen Service
ISO 21011: Cryogenic vessels - Valves for cryogenic service
EIGA 200/17: The Safe Design, Manufacture, Installation, Operation and Maintenance of Valves Used in Liquid Oxygen and Cold Gaseous Oxygen Systems
ISO 23208: Cryogenic vessels – cleanliness for service (specify for LOX service)

## Tanks (General)
Make
Model
Component | Specification | Meet requirement? | Type of supporting documents | Comments
Design Parameters | Vertical construction, stationary tank type
• Stainless steel or carbon steel outer shell
□ Affixed lifting lugs for use w/ clevis pins
• Stainless steel inner vessel (no aluminum)
• Stainless steel piping
• Vacuum pressure 0.05 bar (or better)
Maximum design pressure / test pressure (MAWP): 18 bar
Minimum / maximum operating temperatures:
• Inner vessel: -196°C /+20°C
• Outer shell: -30 °C /+50 °C
Sizing | 20 tonnes
Features | Operating:
• Fill coupling / type for LOX (spec type available)
□ Both bottom and top-fill lines
• Pressure regulation
□ Pressure build-up/blow-down assembly
□ Pressure regulator with set-point feature for MAOP
• Overfill protection system, comprising
□ Fill line drain valve
□ Trycock vent valve
Monitoring:
• Liquid level gauge (& type to be specified)
• Liquid level transmitter (& type to be specified
Safety:
•Thermal relief valve
• Safety valves: primary (set at MAOP) and secondary (set at MAWP) (no bursting discs)
Additional offtake:
• Economizer circuit(preferred option)
• Discharge line for liquid offtake (indicate tanks that this is standard on, and if possible, for smaller units)
Documentation | Detailed description of any relevant experience of importing/suppling liquid medical oxygen storage tanks
Evidence of effectively managed after sales service, including but not limited to installation, testing, and commissioning, repair, warranties honored, and provision of spare parts
P&ID - generic
Sample databook, illustrative of what would be included (not limited to):
• Manufacturer's declaration of conformity
• 3rd party (e.g., Notified Body for PED) certificate of conformity and test reports, certificates for safety valves
• Commissioning, operations, and maintenance manual
• Detailed stress / loading calculations
Export license
Warranty | 2 years minimum
The supplier must ensure the availability of spare parts for the duration of the warranty and indicate lead times
Regulatory | PED 2014/68 EU (indicate NB) Or ASME “U” stamp
QMS | ISO 9001: Quality Management Systems - Requirements
General design guidance (or equivalent) | EN 13458-1: Cryogenic vessels - Static vacuum insulated vessels - Part 1: Fundamental requirements
Process or component specific standards (or equivalent) | EIGA 200/17: The Safe Design, Manufacture, Installation, Operation and Maintenance of Valves Used in Liquid Oxygen and Cold Gaseous Oxygen Systems
ISO 8501-1: Preparation of steel substrates before application of pains and related products – visual assessment of surface cleanliness – Part 1: Rust grades and preparation grades of uncoated steel substrates and of steel substrates after overall removal of previous coatings
ISO 8504-2: Preparation of steel substrates before application of paints and related products – surface preparation methods – part 2: abrasive blast-cleaning
ISO 12944-4: Paints and varnishes – corrosion protection of steel structures by protective paint systems – Part 4: Types of surface and surface preparation
ISO 12944-5: Paints and varnishes – corrosion protection of steel structures by protective paint systems – Part 5: Protective paint systems
ISO 21009-1: Cryogenic vessels – static vacuum-insulated vessels – Part 1: Design, fabrication, inspection, and tests
ISO 21009-2: Cryogenic vessels – Static vacuum insulated vessels – Part 2: Operational requirements
ISO 21010: Cryogenic vessels – Gas/material compatibility
ISO 21013-1: Cryogenic vessels – Gas/material compatibility: Cryogenic vessels – pressure-relief accessories for cryogenic service – Part 1: reclosable pressure relief valves
ISO 23208: Cryogenic vessels – cleanliness for service (specify for LOX service)

## Tanks (Product Specific)
Product availability & associated specifications
Size, Tonnes LOX capacity | 20 tonnes each
Equivalent size, litres
Boil-off rate (indicate conditions)
Discharge capacity
Thermal conductivity
Insulation/fill material
Structural testing: wind & snow loading, earthquake rating

## First fill mLOX manufacturer
INSTRUCTIONS
If you are currently transporting medical LOX, please complete Table 1 (blue) and attach relevant suporting documentation. CHAI wishes to support the entry of new, prepared and quality suppliers to the oxygen market. If you are responding to this bid solicitation as a growth opportunity, please complete Table 2 (green) and attach relevant supporting documentation from existing medical oxygen transportation operations where applicable. Please note, the tables are not evaluation forms with a binary outcome. Data collected will be used to initiate a risk assessment (if deemed necessary) and determine priority for follow-up conversations & evaluations. When a risk assessment is completed, we will move forward with consideration for the current market and regulatory infrastructure in country. If you have not met the requirements in the tables previously because it has not been required by your regulatory and market environment, we wish to understand what preparations you have the capacity to make and where additional technical assistance, resources or time may be required.
Component | Specification | Meets requirement? (Yes/No/NA) | Type of supporting document (select from drop down) | File name as submitted | Page No. | Comments
Volume | 20 tonnes
Regulatory | Proof of stringent regulatory authority (SRA) approval for medical LOX.
OR
Proof of product meeting the requirements of an SRA pharmacopeia (e.g., USP, Ph Eu, BP, or IP) as shown by their specification and a sample Certificate of Analysis (CoA).
Good Manufacturing Practices(GMP) | Address of the manufacturing facility used for the first fill during commissioning
Certificate of Analysis (CoA) template to indicate alignment with USP, Ph Eu, Int’l Pharmacopoeia, or British Pharmacopoeia
Copy of the manufacturer’s license to distribute medical oxygen or license as a wholesaler of medical oxygen
Proof of GMP OR the following:
• Certification of ASU, and storage tanks for oxygen service for medical application.
• Cleaning protocols for tanks (storage and transport) (e.g., cleaned to ISO 23208 with ‘cleaned for O2’ indicated in scope).
• Audit history of a manufacturing facility from the past 3 years (including validity dates for audits).
• Facility safety directives.
• If a manufacturer is producing industrial and medical LOX, an explanation of when the product streams are separated and how contamination is prevented.
• Explanation of how a batch of medical LOX is defined.
• Validation of manufacturing and cleaning processes.
Please provide a letter of authorization from between medical LOX manufacturer and LOX infrastructure supplier if medical LOX supply is outsourced.
Quality Management Systems | ISO 9001: Quality Management Systems - Requirements
Documentation | Evidence of sale of medical LOX to government body
Letter of authorization between mLOX manufacturer and LOX infrastructure supplier if mLOX supply is outsourced.
Please list up to 5 institutional clients (e.g. international NGOs, large corporate bodies, government procurement agencies). Please include pertinent information about the relationship (e.g contract size, duration, products, etc.)

## First fill mLOX transport
INSTRUCTIONS
If you are currently transporting medical LOX, please complete Table 1 (blue) and attach relevant suporting documentation. CHAI wishes to support the entry of new, prepared and quality suppliers to the oxygen market. If you are responding to this bid solicitation as a growth opportunity, please complete Table 2 (green) and attach relevant supporting documentation from existing medical oxygen transportation operations where applicable. Please note, the tables are not evaluation forms with a binary outcome. Data collected will be used to initiate a risk assessment (if deemed necessary) and determine priority for follow-up conversations & evaluations. When a risk assessment is completed, we will move forward with consideration for the current market and regulatory infrastructure in country. If you have not met the requirements in the tables previously because it has not been required by your regulatory and market environment, we wish to understand what preparations you have the capacity to make and where additional technical assistance, resources or time may be required.
Component | Specification | Meets requirement? (Yes/No/NA) | Type of supporting document (select from drop down) | File name as submitted | Page No. | Comments
Volume | Single delivery of 20 tonnes to CUK
Good Manufacturing Practices(GMP) | Address of the transportation company headquarters and any relevant subsidiary for a company that will transport mLOX and carry out first fill procedures.
Copy of the transporter/supplier’s license to distribute medical oxygen or license as a wholesaler of medical oxygen
Copy of export license for the transport of medical oxygen if border crossing
Evidence of relevant certifications (e.g., TPED 2010/35/EU or US DOT) and labeling (e.g., GHS, DOT hazard class, UN ID 1073 class) of road tanker for oxygen service for medical application).
Cleaning protocols for transportable cryogenic vessels (e.g., cleaned to ISO 23208 with ‘cleaned for O2’ indicated in scope).
Safety directives, procedures, or internationally recognized guidelines (e.g., EIGA) that are followed to ensure safe transport and transfilling of medical LOX
If transporter distributes industrial and medical LOX, explanation of how contamination of transport vessels is prevented
Explanation of how a batch of medical LOX is defined.
Validation of cleaning processes.
Evidence of 'first fill’ procedures that will be followed during final testing and commissioning to transfill medical LOX from a road tanker to a newly installed empty bulk storage tank.
Evidence of maintaining pressure differential between tanker truck and bulk storage tank during filling to prevent backflow.
A letter of authorization from the company making the deliveries to the company supplying the LOX (if applicable).
Quality Management Systems | ISO 9001: Quality Management Systems - Requirements
Documentation | Evidence of sale of medical LOX to government body
Please list up to 5 institutional clients (e.g. international NGOs, large corporate bodies, government procurement agencies). Please include pertinent information about the relationship (e.g contract size, duration, products, etc.)
TABLE 2: not currently transporting medical LOX
# | Question | Response |  |  |  | File name(s) of attached supporting documentation (if applicable)
1 | Address of the transportation company
2 | Please provide status or plans to obtain appropriate licenses for transport and distribution of medical oxygen.
3 | Do you plan to transport industrial and medical LOX? If so, please provide an explanation of how industrial and medical LOX will be segregated.
4 | Please describe how medical LOX transport vehicles will be inspected for oxygen service. Include what procedures will be implemented.
5 | Please provide overview of 'first fill' procedures that will be followed during final testing and commissioning. Include what procedures will be implemented.
6 | Please describe medical LOX transportation safety directives you will implement, and whether you will be able to adopt and modify any safety directives from existing industrial LOX or gas cylinder transport operations.
7 | Have you been hired to distribute medical oxygen to a government body? If so, please describe how the medical oxygen was distributed (gas cylinders or cryogenic storage vessel, LOX or gaseous oxygen, etc)
8 | Please list up to 5 institutional clients (e.g. international NGOs, large corporate bodies, government procurement agencies). Please include pertinent information about the relationship (e.g contract size, duration, products, etc.)

## RFI_MFLab (hidden)
| 关于产品 | [9 empty] | 关于贵司
产品类型 Product Type | 产品标准要求 Product Specifications | 产品型号（请列出贵公司符合左侧标准要求的产品，并邮件提供电子版产品用户手册） Product Model (s) | 产品排期（假设下周下订单，预计能够供的最早出货排期至何时（日期）？ Lead time if order placed next week (date) | 本次订单所需数量 Units Required | 公司当前可出售的数量 Units Available | 法规批准/认证Regulatory Approval/Certification |  | 产品性能标准Standards, for product performance |  | 产品单价（美金） Unit price (USD) | 制造商标准 Standards, for manufacturer |  | 预付款要求 Pre-payment requirement | 公司以前有向埃塞俄比亚出口的经验吗？Does company have previous export experience to Ethiopia? | 图片示例 Picture
[6 empty] | 标准 | 如果有，请填写 “X” | 标准 | 如果有，清填写 “X” |  | 标准 | 如果有，清填写 “X”
简单氧气气面罩 Simple Face Masks | 成人式： 医用级 | LB1110, Adult(LA), 5.5mm x 2m绿色管 |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X | 生产前付清 | 有
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
不重吸面罩 Oxygen Mask with reservoir bag | 成人式： 氧气浓度：支持80%至100％的氧气浓度 管长: 2m 使用标准接头；管路（氧气管线）不弯曲 材质：聚氯乙烯（PVC）； 透明；医用级 | LB1310, Adult(LA), 5.5mm x 2m绿色管 氧袋1000ml |  | 2000 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
ISO 18190
[8 empty] | ISO18562-1 | X
ISO/DIS 23368
[8 empty] | ISO/DIS 17256 | X
[8 empty] | ISO 15223-1 | X
文丘里面罩（氧浓度可调面罩） Venturi Mask | 成人式：能以8至10 L / min 的流速输送低至中等流量的氧气，特定氧浓度为30％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,31%(氧气流速6L/min±10%),白色，其他符合 |  | 300 |  | CE | X | ISO1712:2009 |  |  | ISO9001 | X
[6 empty] | FDA | X | ISO15001 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO18562 | X |  | ISO14971 | X
| 成人式：能以10至12 L / min 的流速输送低至中等流量的氧气，特定氧浓度为40％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,40%(氧气流速8L/min±10%),粉色，其他符合 |  | 300 |  | CE | X | ISO 18190
[6 empty] | FDA | X | ISO18562-1 | X
[6 empty] | 出口许可证 | X | ISO/DIS 23368
| 成人式：能以12至15 L / min 的流速输送低至中等流量的氧气，特定氧浓度为60％。 面罩具有可调节的鼻夹。 该套件包括连接管，湿度杯和调节阀等，调节阀用颜色编码并指示氧气的百分比 (最好黄颜色）。 | LB1220,60%(氧气流速8L/min±10%),红色，其他符合 |  | 300 |  | CE | X | ISO/DIS 17256 | X
[6 empty] | FDA | X | ISO 15223-1 | X
[6 empty] | 出口许可证 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 中等大小 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带, M |  | 130 |  | CE |  | ISO 18562-1:2017 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 17510:2015 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X |  |  |  | ISO14971 | X
非侵入性/ CPAP口罩 Non-Invasive/CPAP Masks | 入口细菌过滤器（如果适用）（至少180）。 呼气过滤器，效率高。重复使用，可以承受高级消毒和灭菌。 大码。 | LB3041,不带细菌过滤器，全硅胶，面罩+支架+头带，L |  | 130
气管导管 Endotracheal Tube | 管： 尺码：7。 内径：7.0mm。 外径：9.3mm。 长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 袖口： 直径：24mm | LB5010C, 带囊，ID 7.0mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X |  | ISO9001 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X |  | ISO13485 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X |  | ISO14971 | X
| 管： 尺码：7.5 内径：7.5mm。长度：最小320mm。 接头： 近端：标准内径：15mm。 远端（适合管子）：长度：最小16mm。 | LB5010C, 带囊，ID 7.5mm,其他符合 |  | 350 |  | CE | X | ISO 5361:2016 | X
[6 empty] | FDA |  | ISO 10993-1:2018 | X
[6 empty] | 出口许可证 | X | ISO 11135:2014 | X