## Glucometers and Strips
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
General Description | Portable, battery-powered, semi-automated or automated instrument (self-testing meter) for quantitative measurement of glucose in whole blood specimens.
Intended Use | Handheld device (approx. 6 cm x 12 cm x 3 cm). For diagnosis and management of Diabetes mellitus, hypoglycemia, hyperglycemia in adults and children.
Battery: One set provided; preferred – extra batteries for ≥18 months of use (20 tests/day).
Technical / Performance | Measurement range ≥30–400 mg/dL (1.7–22.2 mmol/L); preferred 20–500 mg/dL (1.1–27.8 mmol/L).
Accuracy: 95% of readings within ±15% for ≥100 mg/dL and ±15 mg/dL for <100 mg/dL; 99% of results within zones A/B of Consensus Error Grid.
Precision (Repeatability): CV <5%. Bias typically <15% (preferably <10%).
Analytical specificity: No interference from ascorbate, lipid, protein, or hematocrit variation.
Limit of Detection: ≤30 mg/dL (1.7 mmol/L), preferred ≤20 mg/dL (1.1 mmol/L).
Operations | Assay principle: Enzymatic (glucose oxidase or glucose dehydrogenase).
Sample volume: <15 µL (preferred <10 µL).
Result display: Quantitative in mmol/L or mg/dL; clear digital screen.
Time to result: <1 minute (preferred <30 seconds).
Monitoring & Alarms | Audible/visual alarms for: low battery, malfunction, strip error, expired strip, and high/low glucose levels.
Environmental Conditions | Operates within 5–45 °C, RH 10–90% non-condensing.
Power Requirements | Internal battery (no memory loss when removed).
If rechargeable: AC 100–240V (±10%), 50–60Hz; full charge ≤3 hours.
Battery life ≥30 days at 20 tests/day (≥600 tests total; preferred ≥900).
Warranty | Minimum 1 year (preferred 2 years).
Quality Management System (QMS) | Manufacturer compliant with ISO 13485 (QMS for medical devices) and ISO 14971 (risk management).-Provide proof of QMS such as ISO 13485 certificate and proof of SRA such as FDA approval, EC certificate, Health canada approval , etc) Letter Of Authorisation (LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Consumables & Accessories | Test strips: Compatible, 20-month shelf life from supply date; ≥2 months after opening.
Control solutions: Low, medium, and high levels, or calibration materials as required.
Product Standards | IEC 60601-1-2, IEC 60601-1-6, ISO 15197

## Lancets
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | Single-use retractable safety lancets for finger pricking. (SOURCE: GMDN 61579, UNICEF S0531996)
Technical | Pre loaded pressure operated (preferable)
Easy handling, two step operation and button operated design (preferable)
Single-use, unable to be reused
Needle is fully shielded before and after use
Needle automatically retracts the needle into casing after sampling
Edged, high quality ultra sharp needle, which can rapidly penetrate skin to reduce pain
Sterile individually packed
Initial sterilisation method: Ethylene oxide gas or Gamma radiation
Typical incision depth, approximately 2mm
Intended use | For use with both adult and paediatric patients
QMS | ISO 13485 (medical device QMS)
Regulatory | Proof of SRA such as CE / FDA,/Health Canada etc Letter Of Authorisation (LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA registration Status | Has valid Ghana FDA registration certificate

## HBA1c Strips
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | Disposable, single-use test strips designed for quantitative measurement of glycated haemoglobin (HbA1c) in whole blood (capillary and/or venous) using a compatible point-of-care analyser. Results are used to monitor long-term glycaemic control and support diabetes diagnosis.
Intended Use | Professional use in healthcare facilities and diabetes clinics for monitoring HbA1c levels in adults and children.
Principle / Methodology | Enzymatic, immunoassay, or boronate affinity method traceable to IFCC reference system and standardized to NGSP/DCCT values.
Sample Type | Capillary (finger-prick) and/or venous whole blood (EDTA).
Sample Volume | ≤ 5 µL whole blood.
Measurement Range | 4.0 % – 15.0 % HbA1c (20–140 mmol/mol, IFCC units).
Accuracy (Bias) | Bias from reference method at clinically relevant HbA1c (~48 mmol/mol / 6.5 %) must not exceed ± 6 % (≈ ± 0.4 % HbA1c) in ≥ 95 % of results.
Precision (CV) | Total coefficient of variation (CV) ≤ 3 %.
Analytical Specificity | Device performance must remain accurate across hematocrit range 25–65 %; demonstrate minimal interference from Hb variants (HbS, HbC, HbE).
Calibration / Traceability | Must be traceable to IFCC reference and NGSP certified.
Time to Result | ≤ 10 minutes per test (preferred ≤ 7 minutes).
Display / Reporting Units | HbA1c reported as % (NGSP) and/or mmol/mol (IFCC), as per manufacturer configuration.
Storage & Stability | Shelf life ≥ 18 months from manufacture; stable under 10–40 °C and ≤ 85 % RH (non-condensing); open-vial stability ≥ 30 days.
Operating Conditions | Ambient 10–40 °C; ≤ 85 % RH (non-condensing).
Packaging & Labelling | Strips individually packaged or in sealed cartridges; clearly labelled with lot number, expiry date, and storage instructions.
Accessories / Inclusions | Compatible analyser, code chip (if required), control/calibration materials, instructions for use.
Quality & Regulatory Compliance | • ISO 13485 (Medical Devices QMS) • ISO 15197 (POCT performance) • IFCC traceability and NGSP certification • CE IVD and/or FDA 510(k) approval Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Warranty / Shelf Life | Shelf life ≥ 18 months; manufacturer warranty ≥ 12 months for analyser compatibility and reagent stability.
Documentation Required | Certificate of analysis, performance validation data, CE/FDA/NGSP documentation, IFU (English).
Ghana FDA registration status | Has valid Ghana FDa registration certificate

## Blood Pressure Monitor
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
General Description | Automated digital blood pressure monitor for upper arm measurement, suitable for adult clinical use. Includes large cuff range and validated for accuracy in hypertensive and obese populations.
Intended Use | For non-invasive measurement of systolic and diastolic blood pressure and pulse rate in adults. Designed for use in hypertension, diabetes, and NCD clinics.
Measurement Method | Oscillometric method, automatic inflation and deflation.
Measurement Range | - Systolic: 60–260 mmHg - Diastolic: 40–200 mmHg - Pulse rate: 40–180 bpm
Measurement Accuracy | - Pressure: ±3 mmHg or ±2% of reading - Pulse rate: ±5% of reading
Display | Large LCD digital display (minimum 20 mm digit height), showing systolic, diastolic, and pulse rate simultaneously.
Memory / Storage | Minimum 90 readings per user (2-user mode preferred).
Power Supply | - AC adapter (100–240 V) and/or AA batteries (≥4 pcs) - Auto power-off to conserve energy.
Cuff Type | - Adult cuff (22–42 cm) included. - Large cuff (32–52 cm) optional or additional for obese patients.
Cuff Material | Durable nylon, washable, latex-free.
Cuff Inflation | Automatic, silent pump with controlled deflation valve.
Operating Conditions | 10–40°C; 15–90% RH non-condensing.
Data Output (Optional) | USB or Bluetooth interface for EMR integration.
Calibration | Factory calibrated to ISO 81060-2 standards; recalibration every 2 years recommended.
Display Indicators | - Low battery warning - Irregular heartbeat detection - Motion error indicator - WHO BP classification bar (optional)
Certifications & Standards Compliance | - ISO 81060-2:2018 (non-invasive sphygmomanometers) - AAMI/ESH/ISO Universal Validation Protocol (2020) - CE Mark (EU MDR 2017/745) - FDA 510(k) (if U.S. approved) - ISO 13485:2016 manufacturer certification
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA | Has valid Ghana FDA registration
Construction | Compact, robust design, easy to clean and transport; designed for high patient throughput.
Accessories | Adult cuff, optional large cuff, AC adapter, 4 AA batteries, and user manual.
Warranty | Minimum 2 years.

## Weighing Scale without Stadiome
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | A high-precision, digital, battery or mains-powered weighing scale designed for measuring adult and adolescent body weight in clinical or outpatient settings such as hypertension or diabetes clinics.
Intended Use | For routine monitoring of adult and adolescent patients’ body weight to support clinical assessment and management of hypertension, diabetes, and related conditions.
Type | Digital (electronic) platform weighing scale — floor-standing type, without height measurement feature.
Measurement Range | Minimum 0–200 kg (preferred up to 250 kg).
Graduation / Readability | ≤ 100 g increments (preferably 50 g).
Accuracy | ±0.1 kg or better throughout the measurement range.
Display | Digital LCD or LED display, clear and backlit, showing weight in kilograms (kg) and optionally in pounds (lb).
Platform Dimensions | Minimum 300 mm × 300 mm (preferably 350 mm × 350 mm) — wide and low profile for patient stability.
Platform Material | Non-slip, easy-to-clean surface (preferably tempered glass, stainless steel, or heavy-duty ABS plastic).
Tare and Hold Functions | Required — should allow zeroing before weighing and maintain displayed weight until recorded.
Power Source | Battery operated (AA/AAA or lithium cells) or dual power (battery + AC adaptor). Auto power-off feature preferred.
Battery Life | Minimum 1 year of typical use (or ≥ 1,000 measurements).
Operating Conditions | Temperature: 10–40°C; Humidity: up to 85% non-condensing.
Load Capacity Protection | Integrated overload protection system.
Weight Units | Kilograms (kg) and pounds (lb), switchable.
Portability | Lightweight, compact, and stable design; ideally < 5 kg for easy movement within clinic.
Quality and Safety Standards | • ISO 13485 (Medical device QMS) • IEC 60601-1 (Electrical safety) • IEC 60601-1-11 (Home/clinical use) • CE mark or FDA approved
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA registration status | Valid Ghana FDA registration
Calibration | Factory calibrated and capable of user calibration; supplied with calibration certificate.
Warranty | Minimum 1 year (preferred 2 years).
Accessories / Inclusions | User manual, calibration certificate, and spare batteries.

## Weighing Scale with Stadiometer
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
General Description | Electronic, hospital-grade weighing scale with integrated mechanical or digital height meter. Designed for use in adult outpatient, hypertension, and NCD clinics.
Intended Use | For measurement of adult body weight and height to support BMI calculation and cardiovascular risk assessment.
Weighing Capacity | Minimum 200 kg (preferably 250 kg) in accordance with OIML R76 and WHO standards.
Accuracy Class | OIML Class III (medical grade).
Graduation (Division Size) | 100 g or better.
Accuracy | ±0.1% of full scale or better.
Platform Size | ≥ 350 × 350 mm, anti-slip, low-profile, stable platform.
Height Measurement Range | 60–210 cm (or wider).
Height Graduation | 1 mm (0.1 cm).
Display | Large digital LCD display (≥25 mm digit height), backlit, visible under clinical lighting.
Tare/Zero Function | Automatic and manual taring/zeroing functions.
BMI Function | Integrated BMI calculation (for digital height meters).
Power Supply | AC power (220–240 V) with rechargeable battery providing ≥40 hours continuous operation.
Battery Type | Rechargeable lithium or lead-acid; auto shut-off to conserve power.
Construction Material | Heavy-duty powder-coated steel or high-grade ABS plastic, corrosion-resistant.
Calibration | Factory calibrated and compliant with OIML R76; capable of field recalibration.
Operating Temperature | 10°C – 40°C, 10–90% RH non-condensing.
Data Output (Optional) | Bluetooth or USB interface for EMR data transfer.
Certifications & Standards Compliance | - OIML R76 (Non-Automatic Weighing Instruments) - EN 45501:2015 (European standard) - ISO 13485:2016 (Quality system) - CE Mark (EU MDR 2017/745) - WHO Recommended Device Specifications
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA registration status | Has valid Ghana FDA regsitration certificate
Accessories | AC adapter, height rod, calibration certificate, and user manual (English).
Warranty | Minimum 2 years.

## Stadiometer
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | A precision instrument designed for measuring the height of standing adults and adolescents in clinical and outpatient settings such as hypertension and diabetes clinics. Intended for use in conjunction with a weighing scale to determine Body Mass Index (BMI).
Intended Use | For accurate and reliable measurement of patient height during routine clinical assessment of adults and adolescents, particularly for BMI calculation and nutritional risk evaluation.
Type | Mechanical or digital stadiometer — wall-mounted, portable, or free-standing type (depending on clinic setup).
Measurement Range | Minimum 20 cm to 210 cm (preferred up to 230 cm).
Graduation / Readability | 1 mm increments (preferred readability ≤ 1 mm).
Accuracy | ±2 mm or better across full measuring range.
Measuring Scale | Clearly graduated in centimeters (cm) and millimeters (mm); dual-scale in inches optional. Permanent, wear-resistant print.
Headpiece | Smooth, flat, right-angled headpiece with easy vertical movement and firm locking mechanism; should rest squarely on the head for accurate readings.
Construction / Material | Durable, non-corrosive materials such as anodized aluminum or high-strength ABS plastic. Surfaces must be easy to clean and disinfect.
Base (for freestanding models) | Wide, stable, non-slip base to ensure patient safety and measurement stability.
Portability (for portable models) | Lightweight design, foldable or sectional with carrying case; total weight not exceeding 5 kg.
Installation (for wall-mounted models) | Mounting bracket and fixings provided; should attach securely to vertical wall surface.
Display (for digital models) | Digital LCD or LED display with clear height readout; battery-powered or dual power (AC/battery).
Operating Conditions | Temperature: 10–40°C; Humidity: up to 85% non-condensing.
Cleaning / Maintenance | Smooth surface, resistant to standard hospital disinfectants; no exposed mechanisms requiring lubrication.
Quality and Compliance Standards | • ISO 13485 (Medical Device QMS) • CE mark or FDA approved • IEC 60601-1 (for digital types, electrical safety)
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA registration status | Has valid Ghana FDA regsitration certificate
Accessories / Inclusions | User manual, installation guide, and calibration certificate (if applicable).
Calibration | Factory calibrated; capable of re-calibration where required.
Warranty | Minimum 1 year (preferred 2 years).

## Measuring Tape
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
General | Children's Mid Upper Arm Circumference (MUAC) measuring tape with cut-off point for severe acute malnutrition at 11.5 cm, pack of 50 tapes with text and pictorial instructions for use printed on the reverse side, reusable, in English language.
Technical Specifications | Tape for measuring child's Mid Upper Arm Circumference (MUAC).
Measuring range up to 26.5 cm.
Graduated with 1 mm precision with thicker line at 21.0 cm.
Accuracy: ± 1 mm of the maximum measurement (26.5 cm).
With three incisions to facilitate tape insertion and fixation, including one wider result window.
Double-sided, featuring user instructions in English (‘tape method’) on the reverse side.
Print: Permanent, resistant to solvents; easily readable in low light working situations.
Working/storage temperature range: 0 - 50 °C.
Material | Non-tear stretch-resistant polypropylene.
Minimum thickness 0.3 mm.
Materials used are free from phthalates, heavy metals and other dangerous matters as per the latest Candidate List of substances of very high concern for Authorisation published on the website of the European Chemical Agency.
Front side | Colour-coded as follows:
Red (Pantone code 1795 C): from 0 to 11.5 cm,
Yellow (Pantone code 107 C) from 11.5 to 12.5 cm,
Green (Pantone code 369 C) from 12.5 to 26.5 cm.
Reverse side | Pictorial and text user instructions in English.
Brand name (not mandatory).
Documentation to be submitted with bid | Supplied with: 1 x Instructions for use leaflet in English

## Urine Dipsticks
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description / Intended Use | Multiparameter reagent strips for urinalysis (≥10 parameters).
Parameters | Glucose, protein, pH, blood, ketone, bilirubin, urobilinogen, nitrite, leukocytes, specific gravity (≥10).
Reading | Visual or compatible with automated urine analyzers.
Result Time | ≤2 minutes.
Storage | 2–30°C; humidity-resistant container with desiccant.
Shelf Life | ≥24 months (≥6 months after opening).
Packaging | 50 or 100 strips per vial.
Standards / Certification | ISO 15197, ISO 13485 certification, and Proof or SRA (EC IVDcertificate, FDA approval, Health Canada approval, etc)
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA registration status | Has valid Ghana FDA regsitration certificate

## Monofilament
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | A standardized nylon monofilament device designed to assess loss of protective sensation in patients, particularly for screening diabetic peripheral neuropathy in adults and children.
Intended Use | For clinical screening of peripheral sensory neuropathy by applying controlled pressure to specific sites on the feet (and sometimes hands) of patients with diabetes or other neuropathic conditions.
Type | Handheld, single-strand monofilament calibrated to deliver approximately 10 grams of force when bent to form a “C” shape.
Material | High-quality nylon filament attached to a firm plastic base or handle; resistant to deformation and environmental humidity.
Calibration Force | 10 grams ± 1 gram at full deflection (as per Semmes–Weinstein standard).
Length of Filament | Approximately 38 mm (standard).
Base / Handle | Ergonomic, lightweight handle to ensure accurate and comfortable application by the clinician.
Reusability | Reusable for a limited number of patients if properly stored and disinfected; replace when filament shows signs of damage, bending, or loss of calibration.
Sterilization / Cleaning | Can be disinfected with mild antiseptic wipes (e.g., 70% isopropyl alcohol); not autoclavable.
Storage Conditions | Store in protective case at room temperature, away from sunlight and moisture to maintain calibration.
Packaging | Individually packed or in sets (e.g., 5 or 10 pieces) with storage case and user instructions.
Accessories / Inclusions | Instruction leaflet with illustrated testing sites and procedures.
Operating Conditions | Room temperature (10–40°C); relative humidity ≤ 85%.
Quality and Compliance Standards | • ISO 13485 (Medical Devices QMS) • CE mark or FDA approved • Calibrated according to Semmes–Weinstein standards.
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Warranty / Shelf Life | Typically 12 months or per manufacturer’s recommendation; replace earlier if filament loses calibration.

## Tuning Fork
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | A medical-grade metal tuning fork designed to produce a constant frequency of 128 Hz for assessing vibration perception in patients, particularly to screen for peripheral neuropathy in diabetes care.
Intended Use | Used for clinical assessment of vibration sense on bony prominences (e.g., great toe, ankle, wrist) to detect loss of protective sensation or early neuropathy in adults and children.
Type / Frequency | 128 Hz vibration frequency (standard for neurological examination).
Material | High-quality aluminum alloy or non-magnetic stainless steel to ensure consistent resonance and durability.
Construction Features | Two precisely machined prongs with a calibrated stem; stem length 75–100 mm for stable placement on test site.
Dimensions | Total length approximately 18–20 cm; weight 80–120 g.
Calibration Accuracy | Frequency accuracy ± 1%.
Optional Features | May include fixed damping weights or graduated scale for quantitative vibration threshold measurement.
Surface Finish | Smooth, corrosion-resistant, and easy to disinfect.
Sterilization / Cleaning | Can be cleaned with 70% isopropyl alcohol or mild disinfectant; not autoclavable.
Storage Conditions | Store in protective pouch or case to prevent bending or surface damage.
Packaging | Individually packaged in a protective case with user instructions.
Accessories / Inclusions | Instruction leaflet describing correct testing technique and interpretation.
Operating Conditions | Ambient temperature (10–40°C); relative humidity ≤ 85%.
Quality and Compliance Standards | • ISO 13485 (Medical Devices QMS) • CE marked or FDA approved • Manufactured to medical diagnostic instrument standards.
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Warranty / Shelf Life | Minimum 12 months manufacturer warranty; replace if bent, corroded, or frequency changes.

## Pulse Oximeter
Component | Specification | Meet requirement? | Type of supporting documents | Comments (e.g pg number/additional comments)
Description | Portable, non-invasive fingertip pulse oximeter designed to measure functional oxygen saturation (SpO₂) and pulse rate in adults and children. Suitable for spot-check monitoring in outpatient, primary care, and hypertension/diabetes clinic settings.
Intended Use | Used to measure peripheral oxygen saturation (SpO₂) and pulse rate in patients with hypertension, diabetes, or other chronic conditions to detect hypoxemia or circulatory abnormalities.
Measurement Parameters | SpO₂ and Pulse Rate (PR).
Measurement Range | SpO₂: 35–100%; Pulse Rate: 25–250 bpm.
Accuracy | SpO₂: ±2% (70–100% range); Pulse Rate: ±2 bpm or ±2% (whichever is greater).
Display Type | OLED or LCD digital display showing SpO₂, pulse rate, pulse bar graph, and plethysmographic waveform (optional).
Sensor Type | Dual-wavelength LED sensor (red and infrared) with photodiode receiver.
Response Time | ≤ 10 seconds for stable reading.
Auto Power Function | Automatic power-on when finger inserted and power-off after 5–10 seconds of inactivity.
Power Supply | 2 × AAA alkaline or rechargeable batteries (1.5V each) or USB rechargeable lithium battery (depending on model).
Battery Life | ≥ 24 hours of continuous operation (with alkaline batteries).
Display Orientation | Multi-directional or automatic rotation (preferred for ease of viewing).
Alarm / Indicator Features | Optional visual or audible low SpO₂ and low battery alarms.
Body Construction | Lightweight, high-impact ABS plastic housing; ergonomic design for adult and pediatric fingers.
Dimensions / Weight | Approx. 60 × 35 × 30 mm; weight ≤ 60 g (including batteries).
Environmental Conditions | Operating temperature: 5–40°C; humidity: 15–95% (non-condensing).
Storage Conditions | 0–60°C; ≤ 95% relative humidity.
Cleaning and Disinfection | Wipe with 70% isopropyl alcohol or hospital-grade disinfectant. Not autoclavable.
Accessories / Inclusions | Carry case or lanyard, 2 × AAA batteries (or USB cable if rechargeable), and user manual.
Quality and Safety Standards | • ISO 80601-2-61 (Particular requirements for pulse oximeter equipment) • IEC 60601-1 (General safety) • ISO 13485 (QMS for medical devices) • SRA Approval e,g CE marked or FDA 510(k) cleared.
Letter Of Authorisation(LOA) between the supplier and the manufacturer if the supplier is different from the manufacturer
Ghana FDA registration status | Has valid Ghana FDA regsitration certificate
Warranty / Service Life | Minimum 1-year manufacturer warranty. Expected life span: ≥ 3 years with proper care.

## Medicines
Product Name | Formulation / Dosage Form | Active Ingredient (Compound / Salt) | WHO Prequalification / SRA Approval | Ghana FDA Registration Status | Shelf Life (Months) | Supporting Document & Page number | Comments
Amlodipine 5 mg | Tablet | Amlodipine besylate equivalent to 5 mg amlodipine
Amlodipine 10 mg | Tablet | Amlodipine besylate equivalent to 10 mg amlodipine
Bendroflumethiazide 2.5 mg | Tablet | Bendroflumethiazide 2.5 mg
Hydrochlorothiazide 12.5 mg | Tablet | Hydrochlorothiazide 12.5 mg
Losartan 50 mg | Tablet | Losartan potassium 50 mg
Losartan 100 mg | Tablet | Losartan potassium 100 mg
Lisinopril 10 mg | Tablet | Lisinopril dihydrate equivalent to 10 mg lisinopril
Lisinopril 20 mg | Tablet | Lisinopril dihydrate equivalent to 20 mg lisinopril
Metformin 500 mg | Tablet | Metformin hydrochloride 500 mg
Glimepiride 1 mg | Tablet | Glimepiride 1 mg
Glimepiride 2 mg | Tablet | Glimepiride 2 mg
Gliclazide 40 mg | Tablet | Gliclazide 40 mg
Gliclazide 80 mg | Tablet | Gliclazide 80 mg
Tolbutamide 500 mg | Tablet | Tolbutamide 500 mg
Nifedipine SR 30 mg | Sustained Release Tablet | Nifedipine 30 mg (modified release)
Nifedipine SR 20 mg | Sustained Release Tablet | Nifedipine 20 mg (modified release)